## Slide 1: Examples transferring methods from the Mastersizer 2000 to the Mastersizer 30000
Dr Stephen Ward-Smith MRSC

## Slide 2
First need to briefly recap the methodology of transfer from last months webinar.
Go through several examples of transfer

## Slide 3: Using risk assessment tools for method transfer

## Slide 4: Risk assessment
For liquid dispersion methods
Measurement
Sampling Method
Weighing Method
Laser Alignment
Optical Properties
Measurement time
Humidly
Temperature
Sample transfer to Instrument
Sample Source / Lot
Verification Materials
Instrument Model
Analysis Range
Cleanliness
Sample quantity
Dispersant Source
Instrument
Analyst
Environment
Method
Material
Laser Obscuration
Sample preparation
Pre-dispersion method
Dispersant Grade
Reference method
Analysis settings
Dispersion state:
Sonication power/time
Pump stir speed

## Slide 5: Risk assessment
For dry dispersion methods
Measurement
Weighing Method
Laser Alignment
Optical Properties
Dispersion Pressure
Feed Rate
Measurement time
Extraction
Humidly
Temperature
Sample transfer to Instrument
Sample Source / Lot
Verification Materials
Instrument Model
Analysis Range
Cleanliness
Air Supply
Sample quantity
Air quality
Instrument
Analyst
Environment
Method
Material
Laser Obscuration
Reference method
Analysis settings

## Slide 6: Instrument model and analysis range
Dynamic range
Mastersizer 2000 | Mastersizer 3000
0.02µm to 2000µm | 0.01µm to 3500µm

## Slide 7: Instrument model and analysis range
Hydro 2000S | Hydro 2000G | Hydro 2000MU | Hydro 2000SM | Hydro 2000mP
120ml | 1000ml | 1000ml | 80-120ml | 18ml
Hydro MV | Hydro LV | Hydro EV | Hydro SM | Hydro SV
120ml | 600ml | 600ml/1000ml | 80-120ml | 6-7ml
Liquid dispersion units

## Slide 8: Instrument model and analysis range
Dry dispersion units
Scirocco |  |  | Scirocco M |  | 
Aero S |  |  | Aero M | Aero Funnel Sample Feeder | 
 | Standard | High energy |  |  | 

## Slide 9: Laser obscuration
Changes in the optical design have reduced the susceptibility of the system to multiple scattering

## Slide 10: Dispersion state Liquid

## Slide 11: Achieving comparable wet dispersion
In-line sonication can reduce required ultrasound duration
  Use an ultrasound titration to achieve the same state of dispersion
Ultrasound titration

## Slide 12: Differences in the sampler settings
Mastersizer 3000 dispersion units have a combined pump and stir

## Slide 13: Stir speed titration
Confirm sample suspension with stirrer speed titration
Chart: Stir speed / rpm
| 1000 | 1500 | 1750 | 2000 | 2250 | 2500 | 2750 | 3000 | 3250 | 3500
d10 | 11.74 | 22.32 | 22.62 | 26.04 | 27.34 | 27.86 | 28.34 | 28.66 | 28.68 | 28.46
dv50 | 29.4 | 58.52 | 57.84 | 64.72 | 67.02 | 68.14 | 68.82 | 68.98 | 68.7 | 68.22
d90 | 71.14 | 137.8 | 136.4 | 149.4 | 153.4 | 155.6 | 156.4 | 155.8 | 155 | 154.2

## Slide 14: Dispersion state dry

## Slide 15: Achieving comparable dry dispersion
Degree of dispersion is controlled by primary air pressure
  Monitor change in size distribution with pressure
    Carry out pressure titration

## Slide 16: Comparing pressure titrations on Scirocco & Aero
Standard and high energy dispersers available
Comparable results may be at different pressures
Chart: Air Pressure / bar
| 4 | 3 | 2 | 1 | 0.5 | 0.1
Aero Standard | 82.5 | 84.5 | 93.8 | 101 | 112 | 133
Aero High Energy | 62.3 | 67 | 74.9 | 83.6 | 92 | 115
Scirocco | 73.443 | 72.722 | 75.63 | 83.955 | 93.721 | 120.408
Wet Dispersion | 84 | 84

## Slide 17: Examples

## Slide 18: Emulsion data as supplied

## Slide 19: Some concerns
Older instrument seeing more fines and coarse than new one (which has more capability)
Optical properties used were 1.475, 0.01
For emulsions imaginary RI is normally in the order of magnitude of 0.001
Residual fell from 0.66% to 0.35%
Look at stir speed next

## Slide 20: Pharmaceutical example -1
Imaginary RI of 0.1 used on MS2000, ideally want to keep same, sample preparation makes no difference. Could change RI, or use 2000 analysis
Latter preferred as minimal revalidation needed

## Slide 21: Pharmaceutical example - 2
Old method. Single measurement. Due to changes in good measurement practice decided to take multiple measurements
High obscuration, short measurement masked dissolution!!
Transfer is a good time to examine the old method

## Slide 22: Pharmaceutical example - 2
Switched to a new dispersant
Fortunately same specification was able to be used
What happened?

## Slide 23: Graphite
Method examined to check dispersion
Slurry made up in water with a little Igepal
No ultrasound required
Anode material for batteries
Chart: 
d (0.1) | 11.178 | 11.175 | 11.169 | 11.17 | 11.159 | 11.17 | 11.132 | 11.124 | 11.123
d (0.5) | 16.246 | 16.241 | 16.232 | 16.236 | 16.218 | 16.239 | 16.197 | 16.187 | 16.187
d (0.9) | 23.572 | 23.568 | 23.555 | 23.568 | 23.542 | 23.581 | 23.52 | 23.511 | 23.518

## Slide 24: Graphite Method Transfer
Measurement settings | Mastersizer 2000 | Mastersizer 3000
Particle refractive index | 2.6 | 2.6
Particle absorption index | 1 | 1
Dispersant refractive index | 1.33 | 1.33
Particle type / Particle shape | Irregular | Non-spherical
Analysis model | General purpose | General purpose
Analysis sensitivity | Normal | Normal
Obscuration range | 5% -15% | 5% - 15%
Measurement duration (red) | 10s | 10s
Measurement duration (blue) | 10s | 5s
Accessory | Hydro MU | Hydro LV
Stir speed | 2500rpm | 2500rpm

## Slide 25: Graphite Final Results

## Slide 26: Lithium Cobalt Oxide
Method examined to check dispersion
Slurry made up in water with a little Igepal
No ultrasound required
Chart: 
d (0.1) | 4.66 | 4.66 | 4.66 | 4.64 | 4.6 | 4.64 | 4.64 | 4.6 | 4.59
d (0.5) | 8.04 | 8.04 | 8.04 | 8 | 8 | 8 | 8 | 8 | 8
d (0.9) | 13.5 | 13.4 | 13.5 | 13.3 | 13.5 | 13.3 | 13.3 | 13.5 | 13.5

## Slide 27: Lithium Cobalt Oxide Method Transfer
Measurement settings | Mastersizer 2000 | Mastersizer 3000
Particle refractive index | 1.74 | 1.74
Particle absorption index | 1 | 1
Dispersant refractive index | 1.33 | 1.33
Particle type / Particle shape | Irregular | Non-spherical
Analysis model | General purpose | General purpose
Analysis sensitivity | Normal | Normal
Low obscuration limit | 5% - 15% | 5% - 15%
Measurement duration (red) | 10s | 10s
Measurement duration (blue) | 10s | 5s
Accessory | Hydro MU | Hydro LV
Stir speed | 2500rpm | 2500rpm

## Slide 28: Lithium Cobalt Oxide final results

## Slide 29: Tungsten Carbide
Here we do see the need to optimise dispersion with ultrasound. Need 3 minutes with an external probe to disperse the sample
Chart: Ultrasound duration/ minutes (external)
| 0 | 1 | 2 | 3 | 4
Dx 10 | 0.501 | 0.255 | 0.233 | 0.224 | 0.219
Dx 50 | 1.83 | 0.534 | 0.498 | 0.499 | 0.495
Dx 90 | 49.2 | 3.76 | 2.67 | 2.23 | 2.52

## Slide 30: Tungsten Carbide Method Transfer
Measurement settings | Mastersizer 2000 | Mastersizer 3000
Particle refractive index | 2.24 | 2.24
Particle absorption index | 2 | 2
Dispersant refractive index | 1.33 | 1.33
Particle type / Particle shape | Irregular | Non-spherical
Analysis model | General purpose | General purpose
Analysis sensitivity | Normal | Normal
Low obscuration limit | 5% - 10% | 5% - 10%
Measurement duration (red) | 10s | 10s
Measurement duration (blue) | 10s | 5s
Accessory | Hydro MU | Hydro LV
Stir speed | 2500rpm | 2500rpm

## Slide 31: Initial optical model needed examination

## Slide 32: Tungsten Carbide final results

## Slide 33: In conclusion
It is easy in most cases to transfer methods from Mastersizer 2000 to Mastersizer 3000
Changes to dispersion might be needed (different ultrasound power)
Mastersizer 2000 analysis available

## Slide 34: Acknowlegements
Thank you to everyone who has given me permission to use data in this presentation